FZ-1/4709/SK/17/BCR Katowice, the 2nd of March 2017

[Concern](https://www.diki.pl/slownik-angielskiego?q=concern): [inquiry](https://www.diki.pl/slownik-angielskiego?q=inquiry)

**I. General information**

1. Buyer: Central Mining Institute

Inquiry is conducted in accordance with the principles of competitiveness.

2. For this request for proposal shall not apply to the Public Procurement Law of 29 January 2004.

**II. Object of order**

Purchase of **a radiometer designed for field measurements of ionizing photon radiation.**

***The offer should include:***

- price excluding and including tax (price should include the cost of packing, transport
and insurance from the Seller to the Buyer),

- discounts (if concern),

- date of payment for invoice,

- technical parameters of the radiometer.

***Technical requirements and*** [***parameters***](https://www.diki.pl/slownik-angielskiego?q=parameters)***:***

- The radiometer must enable the measurement of dose rate at least in the range from 20 nSv/h to 1 Sv/h expressed in units of ambient dose rate equivalent,

- The radiometer must detect ionizing photon radiation at least from 45 keV,

- The radiometer must be not heavier than 5 kg, capable of measuring at temperatures in the range from ‑20 °C to +40 °C and work well during precipitations,

- The radiometer must be battery-powered and capable of continuous measurement for at least one month without the need for battery replacement or recharging batteries,

- The measuring set must be equipped with electronic circuits and software for wireless data transfer to computer enabling, and their recording and analysis. In addition, the radiometer should have an internal memory for storing results registered at least within 24 hours,

- The radiometer must be equipped with a tripod

***Other requirements:***

- Term and conditions of delivery – The Buyer requires delivery **within 8 weeks** from the date of order / agreement on conditions DDP Incoterms 2010 marked the place of execution, ie. **Central Mining Institute, 40-166 Katowice, Plac Gwarków 1, SILESIAN CENTRE FOR ENVIRONMENTAL RADIOACTIVITY, Poland.**

- Terms of payment – The Buyer requires payment date, which will be counted from the date of delivery to the GIG valid invoice.

- Conditions of [guarantee](https://www.diki.pl/slownik-angielskiego?q=guarantee) - The Buyer requires at least **12 months** of [guarantee](https://www.diki.pl/slownik-angielskiego?q=guarantee) from the date of receipt of the order.

**III. Criteria of evaluation of offers and selection of the best offer**

**1**.The criteria of evaluation are:

**a) The price = 95%**

**b) Date of payment for invoice = 5%**

**- until 14 days – 0 point**

**- until 21 days – 3 points**

**- until 30 days – 5 points**

**2**. Buyer deems most advantageous and choose the offer that meets all the requirements specified in the Description of the order.

**3.** The price quoted by the Seller shall not be subject to change during the realisation of the order.

**4.** We ​​inform you that if we accept the offer you will be informed of this fact and will be implemented to prepare the order.

**IV. Place and date for submission of offers**

**1**. The offer must be submitted by the 10/03/2017, For hours 09:00 am. The offer can be submitted electronically, by fax or at the Buyer:

**Główny Instytut Górnictwa,**

**Plac Gwarków 1,**

**40-166 Katowice,**

**Poland,**

**fax: 32 259 22 05**

**e-mail:** **skolinska@gig.eu**

**Commercial contact:**

Monika Wallenburg - tel. (32) 259 25 47- fax: (032) 259 22 05 - e-mail: mwallenburg@gig.eu

Sylwia Kolińska - tel. (32) 259 25 55- fax: (032) 259 22 05 - e-mail**:** **skolinska@gig.eu**

**2**. The Buyer informs that the assessment will be subject to only those items that will affect the Buyer during the period from the date of initiation of this market research to date, in which the deadline of submission of offers.

**3**. The Buyer reserves the right to cancel the inquiry at any time, without giving a reason.

**4**. In the course of examining the content of offer The Buyer reserves the right to ask the Seller  for clarification on the content of the offer and to set a deadline to respond of rejection of the offer.

 *Sincerely
  Head of Commercial Department*

 *M.A. Monika Wallenburg*